## Slide 1: Lachat Flow Injection Analysis of Reservoir Water Samples

## Slide 2
Measuring PO4, NH4, NO3, + NO2 in water
2 reservoirs
5 streams
17 sites total
Monthly sampling April-October
Q1: What is the longitudinal heterogeneity of nitrogen, phosphorus, and chlorophyll-a along a double reservoir continuum?
Q2: What are the drivers of heterogeneity among reservoir continua (i.e., residence time, internal loading)?

## Slide 3: Sample collection and processing
Water samples collected at the surface (~0.1m)
Filtered through 0.45 um Whatman GFF filter
Stored frozen until analysis on Lachat
Samples thawed the day before analysis

## Slide 5: Calibration using standards on Lachat
Calibration standards created the day of analysis
  In ppb PO4, NH4, NO3: 1000, 500, 250, 100, 50, 25, 10, 5, 2.5, 0
    Make 1000, 100, 25, and 5 from primary stock
    Other auto-diluted from 1000
  1000 ppb NO2 standard to check nitrate reduction efficiency
  All standards run 3x, integration peaks corrected if necessary, standard curve determined based on the satisfactory peaks
Check standard at 25 ppb, reagent blank, duplicate, and spike (80 uL added to 4mL of sample) run every ten samples

## Slide 6: QAQC of data post-Lachat run
Peaks and integrations checked at the end of the run
  Integrations edited if needed
  Sample re-run if needed
Within QAQC analysis
  Nitrate reduction efficiency
  Duplicate relative % difference
  Spike % recovery
  % error in check standard 25 ppb (pass = <20% diff)

## Slide 7: Nitrate/Nitrite

## Slide 8: Ammonium

## Slide 9
Phosphate